## Slide 1: PETRA veteran cavities.
Unexpected trouble with them
Stefan Wilke, DESY MHF-e
Unexpected trouble with the PETRA veteran cavities
16th ESLS rf meeting
ALBA, Barcelona, 2012-10-09

## Slide 2: Topics.
Short orientation
7 cell cavity
Coupler protection: Detection of voltage breakdown (spark). =
Broken quartz glass window
Temperature regulation of each cavity. Temperature-Map
Coupler ceramic turns partly to grey
Doorknob, round – angled
Misunderstood vacuum event and blue light at coupler

## Slide 3: Short orientation.
rf section
DESY
DORIS
FLASH
XFEL
PETRA: 2304 m, 6 GeV, 100 mA TopUp, 14 undulators (30 beamlines)
Emittance (hor): ε = 1 nmrad ! User operation since 2009
Different fill patterns : 40, 60, 240, 320, 480, (960) bunches
20 MV,
12 cavities (7 cell),
ca. 1600 kW required,
4 klystrons,
max. 3200 kW

## Slide 4: 7 cell cavity, 500 MHz.
For PETRA III we re-installed 12 of the old 16 installed 7-cell cavities.
Mainly because of cost reasons.
Shunt impedance: ca. 23 MOhm
Band width: ca. 54 kHz
Coupler transmission power: ca. 125 kW
Couplers originally designed for 60 kW,
but already operated at DORIS 2 up to 120 kW.

## Slide 5: Strong need to protect the couplers.
This happens on 2010-11-10 at SL_Cy4:
An almost broken ceramic.
Only noticed by a very high difference temperature of the cooling air of the coupler.
support ceramic

## Slide 6: Sudden voltage breakdown in one cavity.
Beam
current
Watched at different cavities. Here SL_Cy6

## Slide 7: .
Detection of sparks
by looking on fast
change of power in
absorbers. This fast
interlock is done by
hardware.
Additionally there are
light sensors.

## Slide 8: Broken quartz glass window.
2011-05-11 at PE_SR_Cy3
This was never seen before.
Reason unknown.

## Slide 9: Individually regulation of cavity temperature.
variation: only max. ca. 3 degree
At some cavities very strong relation: temp. - vac.

## Slide 10: Temperature maps.
The "best"
temperature
for each
cavities is found
empirically.
It depends on
the number of
bunches.

## Slide 11: Pictures from inside a cavity taken by an endoscope.
similar marks of
HOMs (?) seen at
SR_Cy3.
Perhaps we can see
that at every cavity

## Slide 12: HOMs? painting mystery.
Electric
magnetic
Is this the responsible mode?
We need it two times longitudinally!
f = 1,42 GHz?
pictures from
W. Hillert, Bonn

## Slide 13: Grey parts on coupler ceramic.
2010-07-07, SL_Cy2
We see this and similar
pattern on some other
cavities. Also observed
at DORIS.
The cavities with that
grey ceramics did NOT
trip more often!
SR_Cy2
SL_Cy4

## Slide 14: Change of coupler an plungers at SR_Cy3. 2012-04-16.
After many
trips we
changed
the coupler
and the
plungers.
One of them
looks strange.

## Slide 15: Better matching with different doorknob.
original: round
return loss: -7 dB
recovered: angled, multi steps: return loss: -33 dB
40 Bu 80mA | 60Bu 100mA | 240Bu 100mA
22,0°C | 21,8°C | 17,4°C
40 Bu 80mA | 60Bu 100mA | 240Bu 100mA
16,5°C | 18,7°C | 12,3°C
Ceramic window-DT with old „round“ doorknob
Ceramic window-DT with new „angled“ doorknob)
DT= -5,5°C
DT= -3,1°C
DT= -5,1°C
Test at SR_Cy6:

## Slide 16: vacuum event.

## Slide 17: Blue light.
2012-08-06:
at SL_Cy2
starts at ca. 8 kW
but did not expire
up to ca. 35 kW
2011-05-25:
only at cavity SL_Cy4
ca. 30 kW, tuned
light expired at ca. 38 kW

## Slide 18: Removed coupler from SL_Cy2.
2012-08-07

## Slide 19: The last one.
Thank you!